T H E   B B C   M I C R O C O M P U T E R

IEEE488 INTERFACE USER GUIDE

Designed, developed and manufactured in the United Kingdom by Acorn

Computers Ltd and Intelligent Interfaces Ltd

Written by

Andrew G Ray

Intelligent Interfaces Ltd

January 1983

(C)Copyright Acorn Computers Limited and Intelligent Interfaces Limited 1983

Neither the whole or any part of the information contained in, or the product

described in, this manual may be adapted or reproduced in any material form

except with the prior written approval of Acorn Computers Limited (Acorn

Computers) and Intelligent Interfaces Limited (II).

The product described in this manual and products for use with it, are

subject to continuous development and improvement. All information of a

technical nature and particulars of the product and its use (including the

information and particulars in this manual) are given by Acorn Computers and

II in good faith. However, it is acknowledged that there may be errors or

omissions in this manual. A list of details of amendments or revisions to

this manual can be obtained upon request from Acorn Computers Technical

Enquiries. Acorn Computers welcome comments and suggestions relating to the

product and this manual.

All correspondence should be addressed to:

Technical Enquiries

Acorn Computers Ltd

Fulbourn Road

Cherry Hinton

Cambridge CBI 4JN

All maintenance and service on the product must be carried out by Acorn

Computers authorised dealers. Acorn Computers and II can accept no liability

whatsoever for any loss or damage caused by service or maintenance by

unauthorised personnel. This manual is intended only to assist the reader in

the use of the product, and therefore Acorn Computers and II shall not be

liable for any loss or damage whatsoever arising from the use of any

information or particulars in, or any error or omission in, this manual, or

any incorrect use of the product.

Scanned, OCRed, proofread, corrected and reformatted

11-Jan-2008

J.G.Harston

WARNING: THIS APPLIANCE MUST BE EARTHED

IMPORTANT: The wires in the mains lead for this appliance are coloured in

accordance with the following code:

GREEN AND YELLOW  EARTH

BLUE              NEUTRAL

BROWN             LIVE

If the socket outlet available is not suitable for the plug supplied, the

plug should be cut off and an appropriate plug fitted and wired as described.

The plug which is cut off must be disposed of as it would be a potential

shock hazard if it were to be plugged in with the cut off end of the mains

lead exposed. The moulded plug must be used with the fuse and fuse carrier

firmly in place. The fuse carrier is of the same basic colour* as the

coloured insert in the base of the plug. Different manufacturers' plugs and

fuse carriers are not interchangeable. In the event of loss of the fuse

carrier, the moulded plug MUST NOT be used. Either replace the moulded plug

with another conventional plug wired as described, or obtain a replacement

fuse carrier from an authorised BBC Microcomputer/IEEE488 dealer. In the

event of the fuse blowing it must be replaced, after clearing any faults,

with a 3 Ampere fuse that is ASTA approved to BS1362.

As the colours of the wires may not correspond with the coloured markings

identifying the terminals in your plug, proceed as follows:

The wire which is coloured green and yellow must be connected to the terminal

in the plug which is marked by the letter E, or by the safety symbol -!- or

coloured green, or green and yellow.

The wire which is coloured blue must be connected to the terminal which is

marked with the letter N, or coloured black.

The wire which is coloured brown must be connected to the terminal which is

marked with the letter L, or coloured red.

In the event of the fuse blowing it should be replaced, after clearing any

fault, with a 1 amp fuse that is ASTA approved to BS1362.

* Not necessarily the same shade of that colour.

C O N T E N T S

INTRODUCTION ........................................................ 1

Connecting the Interface to the Computer .......................... 1

The IEEE488 Standard .............................................. 1

Information Covered and Where to Find It in this User Guide ....... 2

1 GETTING GOING ....................................................... 3

Simple Data Transfer .............................................. 3

The IEEE Filing System ............................................ 3

Data - strings of up to 255 characters ............................ 5

Data Output ....................................................... 6

Data Input ........................................................ 6

Checking the State of the Interface ............................... 7

An Example of Data Input and Output ............................... 7

2 MORE ON DATA TRANSFER ..............................................  9

Binary Coded Data ................................................  9

The TRANSFER Command ............................................. 10

3 HOW TO FIND OUT WHAT A DEVICE IS DOING ............................. 11

Serial Poll ...................................................... 11

Parallel Poll .................................................... 12

4 OTHER COMMANDS ..................................................... 13

The Timeout Facility ............................................. 13

Remote Local Operation ........................................... 13

Initialising Devices ............................................. 14

Triggering Simultaneous Operation of Devices ..................... 14

5 PASSING CONTROL .................................................... 15

6 THE IEEE FILING SYSTEM - A Summary ................................. 17

Channels ......................................................... 17

Device Addresses ................................................. 18

Device Dependent Messages (Data) ................................. 18

Glossary of IEEEFS Commands ...................................... 21

7 USING THE IEEE FILING SYSTEM IN ASSEMBLER .......................... 51

8 HARDWARE ........................................................... 55

REFERENCES ........................................................... 56

APPENDICES ........................................................... 57

1 IEEEFS Commands - Minimum Abbreviations ........................ 57

2 IEEEFS Errors .................................................. 58

3 Interface Remote Messages ...................................... 59

4 Interface Specification ........................................ 60

5 Fitting the IEEEFS EPROM ....................................... 61

6 Removal of Terminating Resistor Packs .......................... 63

INDEX ................................................................ 64

INTRODUCTION

Before using your IEEE488 interface check that you have received the

following items in addition to this User Guide

- IEEE488 Interface

- IEEEFS EPROM

- Guarantee Card

- Delta Ribbon Cable (for connecting one IEEE488 device to the computer)

If any item is missing, contact the dealer who supplied the IEEE488

interface quoting the order number given to you when you placed your order.

It is strongly recommended that the IEEEFS EPROM be fitted by a dealer.

However, fitting instructions are included in Appendix 5, and information on

the use of the terminating resistors in the IEEE488 Interface are contained

in Appendix 6.

Connecting the Interface to the Computer

1	Turn off the power to the computer and the interface.

2	Plug the 34-way IDC socket on the end of the ribbon cable into the

1MHz Bus header on the underside of the computer. Ensure that pin 1 of

the socket is aligned with pin 1 of the header. The small indents on

the socket should face the printed panel.

3	Turn on the power to the interface and computer.

The IEEE488 Standard

The IEEE488 interface is a general purpose system for exchanging digital

data bit parallel byte serial form between a number of devices in a local

area. The original concept was developed by Hewlett Packard who called it

the HP-IB Hewlett Packard Interface Bus. The IEEE (Institute of Electrical

and Electronics Engineers) first produced a standard in 1975, IEEE Std 488-

1975. This was revised in 1978 as IEEE Std 488-1978. The interface is

identical to that described in the IEC (International Electrotechnical

Commission) Publication 625-1, apart from the mechanical connection system.

The IEEE standard specifies a 24 pin Microribbon connector whereas the IEC

Publication specifies a 25 pin D-type connector as used for the EIA RS232-C

data communication interface.

- 1 -

The IEEE499 interface makes use of two types of messages:

i) interface messages (commands) used to manage the interface

ii) device dependent messages (data)

A device can have the ability to send data (act as a talker), receive data

(act as a listener), or do both (act as a talker-listener), or manage the

system (act as a controller).

This User Guide does not attempt to reproduce the IEEE standard or explain

every aspect of it. The interested user is referred to Ref 1. However, this

User Guide provides all the information required to use the BBC Micro-

computer and IEEE488 Interface with any other IEEE488 device.

Information covered and where to find it in this User Guide

Chapter 1 tells you how to connect the BBC Microcomputer's IEEE488 Interface

to other IEEE488 devices and communicate with them.

Chapter 2 expands on this and covers the exchange of binary coded data.

Some IEEE488 devices have the ability to provide information on what state

they are in, and ways of obtaining this information are discussed in Chapter

3.

Chapter 4 contains details of commands which can be used to change the state

of IEEE488 devices.

The computer can pass control of the IEEE488 interface to another device and

Chapter 5 tells how this is done.

Chapter 6 is useful for reference as it provides a glossary of IEEE488

commands in alphabetical order.

Information on using the interface in Assembler programs is contained in

Chapter 7.

Finally, Chapter 8 gives some details of the hardware used in the interface.

Throughout this User Guide examples are given in BASIC. However, the

interface can be used with any language running on the computer which

permits the use of a filing system. IEEEFS commands are given in full, and

in upper case. However, abbreviations and any mix of lower and upper case

may be used. See Appendix 1 for minimum abbreviations.

- 2 -

1 GETTING GOING - Simple data transfer

Before connecting devices to the computer turn off the power to both

interface and computer.

The interface can be connected to up to 14 other devices using IEEE488

standard cable assemblies. These have a plug and socket at each end,

permitting star or linear interconnection of devices. The connectors are

provided with two securing screws which allow them to be stacked on the

socket of each device. The IEEE socket is normally found on the rear of a

device. The IEEE488 standard permits individual cable lengths of up to 4

metres and a total cable length in a system of 2 metres per device or 20

metres whichever is the shorter. Cables can be obtained from the

manufacturers of devices.

When all devices have been connected to the interface, turn on the power in

the following sequence:

1 devices

2 interface

3 computer

The IEEE Filing System (IEEEFS)

The IEEE488 interface of the BBC Microcomputer appears to the user as an

additional filing system. It gives the computer the ability to act as a

controller-talker or controller-listener with the ability to pass control

and request it back through the use of IEEEFS commands. These commands

enable the full use of an IEEE488 system without detailed knowledge of the

operation of the interface.

The IEEEFS is selected by the command

*IEEE

To reselect another filing system use the appropriate command

*DISK

*TAPE etc

Note, when used with BASIC the IEEEFS does not respond to

*CAT

CHAIN

EOF#

EXT#

LOAD

PTR#

SAVE

- 3 -

Communication between a language and the interface is via two channels. The

first is the 'COMMAND' channel used for transmitting IEEEFS commands and

receiving information on the state of the interface. The second is the

'DATA' channel used for reading and writing data to other IEEE488 devices.

After the first use of *IEEE the COMMAND and DATA channels must be opened.

Example

10 *IEEE

20 cmd%=OPENIN("COMMAND")

30 data%=OPENIN("DATA")

.

.

If required these channels can be closed in the following way

1010 CLOSE#data%

1020 CLOSE#cmd%

Each device in an IEEE488 system must have a unique address. Some devices

have only one address (a primary address) whilst others have extended

addressing (primary and secondary addresses). Primary and secondary

addresses may be in the range 0 to 30. Secondary addresses are often used to

select different functions within the same device, eg an analogue to digital

converter with a number of inputs may have the input selected by the

secondary address. Most devices are delivered with factory set primary and

secondary addresses. These can be found in the manufacturer's manual. If

more than one device in a system has the same address, these can be changed

by following the manufacturer's instructions. Usually this involves changing

switch settings on the rear of the device.

The OPENIN and OPENOUT functions are used to pass the primary and secondary,

if specified, addresses of a device to the IEEEFS via a 'device' channel. Up

to 14 channels may be open at one time.

Example

A digital voltmeter has a primary address of 4 and no secondary address.

100 dvm%=OPENIN("4")

A sixteen input analogue to digital converter has a primary address of 5 and

a secondary address of 11 indicating the input to be converted.

150 adc%=OPENIN("5,11")

If no channel is available a zero is returned.

All channels remain open, even if other filing systems are selected, until

specifically closed while the IEEEFS is active.

A 'device' channel may be closed using the CLOSE# statement.

- 4 -

Example

1110 CLOSE#dvm%

The CLOSE#0 statement closes all 'device' channels but not the COMMAND and

DATA channels which must be closed individually.

As only 14 'device' channels may be open at one time, it is recommended that

channels are closed when no longer required or at the end of a program.

The BBC Microcomputer must also have a unique device address. This is

specified using the BBC DEVICE NO command and can be any number in the range

0 to 30 not already assigned to a device in the system.

Example

To specify the primary address of the computer to 7

40 PRINT#cmd%,"BBC DEVICE NO",7

The CLEAR command is now used to initialise the system.

Example

50 PRINT#cmd%,"CLEAR"

The REMOTE ENABLE command is used to Prepare devices for remote operation,

eg under computer control

Example

60 PRINT#cmd%,"REMOTE ENABLE"

Data strings of up to 255 characters

When strings of ASCII coded data are sent to or received from a device

delimiting characters are used. Consult the manufacturer's manual for each

device to find out the delimiting characters for that device. The computer

defaults to the delimiting character Line Feed (LF, ASCII code 10). If a

device requires different delimiting characters these may be specified using

the END OF STRING command.

Example

To set delimiting characters to carriage return, line feed

80 PRINT#cmd%,"END OF STRING",CHR$(13)+CHR$(10)

The computer will then send these characters and expect to receive them

after each data string until they are respecified.

- 5 -

Data Output

Channels must be opened to the devices concerned as previously described.

Before data can be sent to the devices, they must be listened, using the

LISTEN command.

Example

.

.

160 voltmeter%=OPENIN("9,3")

170 printer%=OPENIN("12")

180 PRINT#cmd%,"LISTEN",voltmeter%,printer%,"EXECUTE"

The EXECUTE command must follow the last device to be listened. Data can

then be sent to the devices through the DATA channel using the PRINT#

statement.

Example

190 PRINT#data%,"R1"

would set the range of the voltmeter and print R1 on the IEEE printer. When

output has been completed, the devices can be unlistened.

Example

340 PRINT#cmd%,"UNLISTEN"

Data Input

Channels must be opened to the devices concerned as previously described.

Before a device can send data it must be talked using the TALK command.

Example

.

.

210 counter%=OPENIN("6")

220 PRINT#cmd%,"TALK",counter%

Note. If there is already an active talker it will automatically be untalked.

Data is then received from the device via the DATA channel using the INPUT#

statement. All active listeners will receive the data sent.

Example

230 INPUT#data%,count$

When input has been completed the device can be untalked.

Example

580 PRINT#cmd%,"UNTALK"

- 6 -

Checking the State of the Interface

The IEEEFS maintains a status word. This is available to the user through the

STATUS command. The meaning of each bit of the status word is detailed in

Chapter 6 under STATUS.

When ASCII coded data is being handled there are three status bits of interest.

bit 4 if set indicates the input sequence was terminated because over 255

characters were received before the delimiting characters

bit 6 if set indicates the END message was received from a device with the

last character

bit 7 if set indicates an incomplete source handshake. This is most likely

to occur during a TALK or LISTEN sequence if no devices are connected

to the interface

Example

.

.

1030 PRINT#cmd%,"STATUS"

1040 INPUT#cmd%,state%

1050 IF (state% AND 64) = 64 THEN PRINT"END received" ELSE PRINT"NO END

received"

.

.

An Example of Data Input and Output

It is required to take three readings in the range 0 to 5v from a digital

voltmeter with a primary address of 9 and print them on the computer's

screen. The voltmeter sends the readings in the format +/- N.NNN CR LF

10 *IEEE

15 DIM voltage$(3)

20 cmd%=OPENIN("COMMAND")

30 data%=OPENIN("DATA")

40 PRINT#cmd%,"BBC DEVICE NO",0

50 PRINT#cmd%,"CLEAR"

60 PRINT#cmd%,"REMOTE ENABLE"

70 PRINT#cmd%,"END OF STRING",CHR$(13)+CHR$(10):REM sets delimiting

characters to CR LF

80 voltmeter%=OPENIN("9")

90 PRINT#cmd%,"UNLISTEN"

100 PRINT#cmd%,"LISTEN",voltmeter%,"EXECUTE"

110 PRINT#data%,"V5"

120 PRINT#cmd%,"UNLISTEN"

130 PRINT#cmd%,"TALK",voltmeter%

140 FOR I% = 1 TO 3

150 INPUT#cmd%,voltage$(I%)

160 PRINT voltage$(l%) :REM display on screen

170 NEXT I%

180 PRINT#cmd%,"UNTALK"

190 CLOSE#voltmeter%

- 7 -

Note In the above example the voltmeter had the same talk and listen

addresses. This is usually, but not always, the case.

Detailed information on each command can be found in Chapter 6, and minimum

abbreviations are in Appendix 1.

- 8 -

2 MORE ON DATA TRANSFER

Binary Coded Data

Sometimes it is necessary to send and receive more than 255 characters at a

time without delimiting characters, so the PRINT# and INPUT# statements

cannot be used.

Devices such as spectrum analysers, digital storage oscilloscopes and other

microprocessor based devices with their own internal memory send and receive

long fixed length sequences of binary coded data. The last byte is normally

accompanied by the END end remote message. Data of this type can be sent and

received using the WRITE BINARY and READ BINARY commands, together with the

BPUT# and BGET# statements.

Example

To send 1024 data bytes held in an array to the memory of a spectrum analyser, primary address 11.

.

.

870 analyser%=OPENIN("11")

880 PRINT#cmd%,"LISTEN",analyser%,"EXECUTE"

890 PRINT#data%,"WTM",4000,1024" :REM prepare analyser to receive data

and store it in memory starting at location &4000

900 PRINT#cmd%,"WRITE BINARY",1024 :REM send the next 1024 bytes and

accompany the last byte with the END message

910 FOR I% = 1 TO 1024

920 BPUT#data%,wave%(I%) :REM output byte

930 NEXT I%

940 PRINT#cmd%,"UNLISTEN"

950 CLOSE#analyser%

.

.

Example

To receive 512 data bytes from the memory of a digital storage oscilloscope, primary address 15, and store them in an array.

.

.

670 scope%=OPENIN("15")

680 PRINT#cmd%,"LISTEN",scope%,"EXECUTE"

690 PRINT#data%,"CH1" :REM prepare oscilloscope to send data

700 PRINT#cmd%,"UNLISTEN"

710 PRINT#cmd%,"TALK",scope%

720 PRINT#cmd%,"READ BINARY",512 :REM prepare to receive 512 bytes

730 FOR I% = 1 TO 512

740 BGET#data%,trace%(I%) :REM input byte

750 NEXT I%

760 PRINT#cmd%,"UNTALK"

770 CLOSE#scope%

- 9 -

The TRANSFER Command

If it is required to exchange ASCII or binary coded data between devices without storing it in the memory of the computer, the TRANSFER command can be used. Two options are available.

a) TRANSFER with a count

Example

To print ten successive readings from a digital voltmeter, primary address 3,

on an IEEE488 printer, primary address 9. The voltmeter sends the readings in

the format +/- N.NNN CR LF ie 8 bytes per reading giving a count of 80 bytes.

.

.

540 printer%=OPENIN("9")

550 voltmeter%=OPENIN("3")

560 PRINT#cmd%,"LISTEN",printer%,"EXECUTE"

570 PRINT#cmd%,"TALK",voltmeter%

580 PRINT#cmd%,"TRANSFER",80

590 PRINT#cmd%,"UNTALK"

600 PRINT#cmd%,"UNLISTEN"

610 CLOSE#voltmeter%

620 CLOSE#printer%

.

.

b) TRANSFER without a count

This can only be used when the END message is sent with the last data byte to

be exchanged. It has the advantage of leaving the computer free to continue

with other asks.

Example

To send a wave form stored in the memory of a fast data capture unit, primary

address 5, to the memory of a spectrum analyser, primary address 8, for

analysis.

.

.

460 analyser%=OPENIN("8")

470 captureunit%=OPENIN("5")

480 PRINT#cmd%,"LISTEN",captureunit%,"EXECUTE"

490 PRINT#data%,"OUT",3000,1024" :REM prepare capture unit to output data

500 PRINT#cmd%,"UNLISTEN"

510 PRINT#cmd%,"LISTEN",analyser%,"EXECUTE"

520 PRINT#data%,"WTM",2000,1024" :REM prepare analyser in input data

530 PRINT#cmd%,"TALK",captureunit%

540 PRINT#cmd%,"TRANSFER",0 :REM transfer with no count

.

REM computer continuing with other tasks

.

910 REPEAT :REM read the status word until the END message is received

920 PRINT#cmd%,"STATUS" :REM causes transfer in progress flag to be reset

930 INPUT#cmd%,state%

940 UNTIL (state% AND 8) = 8

.

.

- 10 -

3 HOW TO FIND OUT WHAT A DEVICE IS DOING

Serial Poll

If a device requires action from the computer, for example in order to

rectify some error condition, it requests this by sending the SRQ (service

request) message. When this happens bit 5 of the status word is set.

The SERIAL POLL command is used to determine which device has requested

service. Each device with a serial poll capability must be polled. Consult

the manufacturer's manual for each device to find out how many bytes it

returns. Most devices return a single byte. Bit 6 of the first byte returned

conveys the RQS (request service) message. If set, this indicates that the

device was requesting service. The remaining bits of the first byte and

ensuing bytes, if any, indicate the reason that service was requested. The

meaning of this information can be found in the manufacturer's manual, and

the appropriate action should be taken.

Example

.

.

270 PRINT#cmd%,"STATUS" 280 INPUT#cmd%,state%

290 IF (state% AND 32) = 32 THEN GOTO 5000

.

.

.

5000 counter%=OPENIN("6'")

5010 signalgenerator%=OPENIN("13")

5020 PRINT#cmd%,"SERIAL POLL",counter%,1 :REM 1 byte expected

5030 INPUT#cmd%,counterstatus$

5040 IF (ASC(counterstatus$) AND 64) = 64 THEN PROCservicecounter

5050 PRINT#cmd%,"SERIAL POLL",signalgenerator%,1

5060 INPUT#cmd%,signalgeneratorstatus$

5070 IF (ASC(signalgeneratorstatus$) AND 64) = 64 THEN PROCservicesignalgenerator

Because each device must be polled separately, a serial poll is time consuming.

- 11 -

Parallel Poll

This is a rapid way of finding out whether one or more devices require

action by the computer. However the computer must be committed to carrying

out a parallel poll periodically.

The PARALLEL POLL REQUEST command returns a single parallel poll response

byte, each bit of which is assigned to a device or devices and, if set,

indicates that the device requires action by the computer.

Devices with a parallel poll capability are assigned a bit in the parallel

poll response byte using the PARALLEL POLL ENABLE command. This command uses

an integer, the least significant 3 bits of which specify the associated bit

of the parallel poll response byte as follows

b2 b1 b0   parallel poll response byte

0  0  0     0  0  0  0  0  0  0  1

0  0  1     0  0  0  0  0  0  1  0

0  1  0     0  0  0  0  0  1  0  0

0  1  1     0  0  0  0  1  0  0  0

1  0  0     0  0  0  1  0  0  0  0

1  0  1     0  0  1  0  0  0  0  0

1  1  0     0  1  0  0  0  0  0  0

1  1  1     1  0  0  0  0  0  0  0

Bit 3 of the integer specifies the 'sense' of the response. Refer to the

manufacturer's manual to decide on its setting.

During a parallel poll request each enabled device compares its individual

status with the value of the sense bit it received when it was enabled. If

the comparison succeeds, the bit assigned to that device is set in the

parallel poll response byte.

If a device is no longer required to take part in a parallel poll, the

PARALLEL POLL DISABLE command is used. The PARALLEL POLL UNCONFIGURE command

renders all devices unable to take part in a parallel poll.

Example

640 siggen%=OPENIN("5")

650 counter%=OPENIN("6")

660 analyser%=OPENIN("10")

670 PRINT#cmd%,"PARALLEL POLL UNCONFIGURE" :REM to ensure no other devices take part in parallel poll

680 PRINT#cmd%,"PARALLEL POLL ENABLE",siggen%,8 :REM sense bit=1 ppr bit0

690 PRINT#cmd%,"PARALLEL POLL ENABLE",counter%,9 :REM sense bit=1 ppr bit1

700 PRINT#cmd%,"PARALLEL POLL ENABLE",analyser%,2 :REM sense bit=0 ppr bit2

710 PRINT#cmd%,"PARALLEL POLL REQUEST"

720 INPUT#cmd%,responsebyte%

730 IF (responsebyte% AND 1) = 1 THEN PROCservicesiggen

740 IF (responsebyte% AND 2) = 2 THEN PROCservicecounter

750 IF (responsebyte% AND 4) = 4 THEN PROCserviceanalyser

.

.

1020 PRINT#cmd%,"PARALLEL POLL DISABLE",analyser%,"EXECUTE" :REM analyser no longer required to take part in parallel poll

.

- 12 -

4 OTHER COMMANDS

The Timeout Facility

The TIMEOUT ON command can be used to enable a 2.5 sec timeout. This is

useful in preventing a program from 'hanging up' when a device fails to

respond. Bit 0 in the status word is set if a timeout occurs during output

and bit 1 is set if a timeout occurs during input. If timeouts are no longer

required the TIMEOUT OFF command is used. Timeouts are not part of the

IEEE488 standard and so the IEEEFS defaults to TIMEOUT OFF.

Example

.

.

500 PRINT#cmd%,"TIMEOUT ON"

510 adc%=OPENIN("17,3")

520 PRINT#cmd%,"TALK",adc%

530 INPUT#cmd%,reading$

540 PRINT#cmd%,"STATUS"

550 INPUT#cmd%,state%

560 IF (state% AND 2) = 2 THEN PROCtimeoutoninput

.

.

.

730 PRINT#cmd%,"TIMEOUT OFF"

.

.

Remote Local Operation

When a device is listened after a REMOTE ENABLE command control of the

device passes from its front panel to the computer. The LOCAL LOCKOUT

command prevents devices with this capability from inadvertently being

returned to front panel control. If it is wished to return control to the

front panel the GO TO LOCAL command is used. However, when the device is

next listened, the front panel switches will again be locked out. To disable

the locking out of the switches, the REMOTE DISABLE command should be used.

This also has the effect of returning all devices to local control.

Example

110 PRINT#cmd%,"REMOTE ENABLE"

120 voltmeter%=OPENIN("21")

130 PRINT#cmd%,"LOCAL LOCKOUT"

140 PRINT#cmd%,"LISTEN",voltmeter%,"EXECUTE" :REM voltmeter under

control with front panel switches ineffective

150 PRINT#data%,"R2" :REM set up voltmeter

160 PRINT#cmd%,"UNLISTEN"

170 PRINT#cmd%,"TALK",voltmeter%

180 INPUT#data%,reading$

190 PRINT#cmd%,"UNTALK"

200 PRINT#cmd%,"GO TO LOCAL",voltmeter%,"EXECUTE" :REM voltmeter under

front panel control

.

.

460 PRINT#cmd%,"REMOTE DISABLE" :REM returns all devices to local

control and disables front panel lockout

470 CLOSE#voltmeter%

- 13 -

Initialising Devices

A device may have been set up incorrectly or an error may have occurred.

Some devices have the capability of being returned to a known condition. The

SELECTED DEVICE CLEAR command may be used to return individual devices to a

known state or the DEVICE CLEAR command to return all devices with the

capability to a known state.

Example

.

.

410 voltmeter%=OPENIN("6")

.

.

.

530 PRINT#cmd%,"SELECTED DEVICE CLEAR",voltmeter%,"EXECUTE" :REM return

voltmeter to known condition

.

.

.

790 PRINT#cmd%,"DEVICE CLEAR" :REM returns all devices with the

capability to a known condition

Triggering Simultaneous Operation of Devices

Sometimes, for example in test systems, it may be required to start a

process and commence measurements at the same time. This can be achieved

using the TRIGGER command which sends the GET (group execute trigger)

message.

Example

660 signalgenerator%=OPENIN("IO")

670 recorder%=OPENIN("9")

930 PRINT#cmd%,"TRIGGER",signalgenerator%,recorder%,"EXECUTE" :REM start

measurement process

- 14 -

5 PASSING CONTROL

The system controller is the device which can send the IFC (interface clear)

and REN (remote enable) messages, and only one is allowed in a system. The

IEEEFS and the interface hardware is supplied configured so that the

computer acts as system controller.

The controller in charge is the device which manages the system. This is not

necessarily the system controller.

After the CLEAR command has been executed the computer is controller in

charge. The TAKE CONTROL command can be used to pass management of the

system to another device.

Example

To pass control to another computer with an IEEE488 interface, primary

address 13

.

.

340 secondcontroller%=OPENIN("13")

350 PRINT#cmd%,"TAKE CONTROL",secondcontroller%

.

.

After control has been passed to another device there are three ways in

which the computer can regain management of the system.

i) When the current controller in charge addresses the computer and sends

the TCT (take control) message.

ii) The REQUEST CONTROL command can be used to indicate to the current

controller in charge that the BBC Microcomputer wishes to regain

control of the system. It does this by sending the SRQ (service

request) message and when polled by the current controller in charge

during a serial poll sends the RQS message indicating it was requesting

service together with a STB (status byte) of zero showing it wishes to

regain control. There is no obligation on the current controller in

charge to pass back control to the BBC Microcomputer. However, if it

does, this occurs as described in the previous paragraph. To find out

when the computer has regained control, bit 2 of the status word should

be checked.

- 15 -

Example

.

.

870 PRINT#cmd%,"REQUEST CONTROL"

880 REPEAT

890 PRINT#cmd%,"STATUS"

900 INPUT#cmd%,state%

910 UNTIL (state% AND 4) = 4

.

.

iii) through the use of the CLEAR command the BBC Microcomputer sends the

IFC (interface clear) message. This instantly halts any current

activity and returns the system to a known state. Data being

transferred may be lost during this process. The computer is then

controller in charge.

Example

1040 PRINT#cmd%,"CLEAR"

- 16 -

6 THE IEEE FILING SYSTEM - A Summary

The IEEE488 Interface appears to the Operating System of the BBC

Microcomputer as an additional filing system. It is selected by the command

*IEEE

The IEEEFS and the interface hardware is supplied so that the computer is

system controller. The computer can act as a controller-talker controller-

listener with the ability to pass control to another device and request it

back.

Channels

Communication between a language and the IEEEFS is via two special channels.

The first is a 'COMMAND' channel used for transmitting commands and

receiving information on the state of the interface. The second is the

'DATA' channel used for reading and writing data.

Example

200 cmd%=OPENIN("COMMAND")

210 data%=OPENIN("DATA")

The IEEEFS maintains a status word, accessible to the user, which enables

recovery from bus communication errors. When a command is received from the

language via the COMMAND channel the IEEEFS interprets it and checks the

status word. If it is satisfactory then the appropriate action is taken and,

finally, the status word is updated. The BBC DEVICE NO and CLEAR command

should always follow the opening of the COMMAND and DATA channels.

- 17 -

Device Addresses

The OPENIN or OPENOUT functions are used to pass the primary and secondary,

if specified, addresses of a device to the IEEEFS via a 'device' channel.

Addresses must be in the range 0-30. Note that the primary address of the

BBC Microcomputer is specified by the BBC DEVICE NO command.

Example

A voltmeter has a primary address of 4 and a secondary address of 2.

100 voltmeter%=OPENIN("4,2")

An analyser has a primary address of 11.

100 analyser%=OPENIN("11")

A 'device' channel may be closed using the CLOSE# statement.

Example

100 CLOSE#voltmeter%

100 CLOSE#0			closes all 'device' channels but not the

'COMMAND' and 'DATA' channels which must be

individually closed.

Device Dependent Remote Messages (Data)

Strings - Input

Strings of up to 255 characters can be input by the IEEEFS and passed to the

language via the 'DATA' channel. The INPUT# statement is used. The IEEEFS

expects to receive the end of string characters after each string input from

device. However, if the END end remote message is received before the end of

string characters this terminates the input sequence.

- 18 -

Remote Messages

ATN

true				computer addresses itself as a listener

false	DAB 1

false	DAB 2

etc

false	DAB n

false	EOS 1			end of string

false	EOS 2 and END	characters

true				computer unaddresses itself as a listener

Example

100 PRINT#cmd%,"END OF STRING",CHR$(13)+CHR$(10)

110 PRINT#cmd%,"TALK",analyser%

120 INPUT#data%,response$

Strings - Output

Strings of up to 255 characters passed to the IEEEFS via the 'DATA' channel

can be output. The PRINT# statement is used. The IEEEFS sends the UNT untalk

remote message before the string and the end of string characters after each

string it outputs to a device. It sends the END end remote message with the

last character.

Remote messages

ATN

true	UNT			computer addresses itself as a talker

false	DAB 1

false	DAB 2

etc

false	DAB n

false	EOS 1			end of string

false	EOS 2 and END	characters

true				computer unaddresses itself as a talker

Example

100 PRINT#cmd%,"END OF STRING",CHR$(13)+CHR$(10)

110 PRINT#cmd%,"LISTEN",analyser%,"EXECUTE"

120 PRINT#data%,"AV4 RE"

- 19 -

Binary Coded Data

If this is to be handled, the READ BINARY and WRITE BINARY commands are

used. Single data bytes are passed to and from the language via the 'DATA'

channel. The BPUT# and BGET# statements are used. For further details see

the READ BINARY and WRITE BINARY commands.

- 20 -

IEEEFS Commands

The following pages describe the IEEE commands under a number of headings.

Purpose

A plain English description of what the command does.

Example

An example to illustrate the use of the command.

Notes

Lists restraints regarding type and range of variables.

IEEE Status Required before Execution

Lists the status flags checked by the IEEEFS before carrying out a command.

IEEE Status after Execution

Lists the status flags changed by the IEEEFS after successfully carrying out

a command.

Remote messages

Shows the logical state of the ATN remote message and the other remote

messages sent by the IEEEFS.

Description

A brief description of the command and its effect.

Associated Commands

Similar commands and those which are concerned with the IEEE488 function.

- 21 -

BBC DEVICE NO

Purpose

To specify the primary address of the computer.

Example

100 *IEEE

110 cmd%=OPENIN("COMMAND")

120 data%=OPENIN("DATA")

etc

200 PRINT#cmd%,"BBC DEVICE NO",6

Notes

Integer constants and variables in the range 0 - 30 may be used for the

primary address.

IEEE Status Required before Execution

No requirement.

IEEE Status after Execution

The	BBC DEVICE NO command executed flag is set.

Remote Messages

None

Description

The computer stores the primary address which is set by the CLEAR command.

It requires a primary address so that, after control has been passed to

another device using the TAKE CONTROL command, it may be addressed as a

talker and control passed back to it. This normally follows a REQUEST

CONTROL command.

Associated Commands

CLEAR

- 22 -

CLEAR

Purpose

To clear the interface.

Example

100 PRINT#cmd%,"BBC DEVICE NO",address%

110 PRINT#cmd%,"CLEAR"

IEEE Status Required before Execution

The	BBC DEVICE NO command executed flag must be set.

IEEE Status after Execution

The	timeout on output flag is reset

timeout on input flag is reset

computer is controller in charge flag is set

TRANSFER in progress flag is reset

string of more than 255 characters received flag is reset

SRQ service request remote message received flag is reset

END end remote message received flag is reset

incomplete source handshake flag is reset

binary coded data being handled flag is reset

a device has been TALKed flag is reset

a device has been LISTENed flag is reset

CLEAR command executed flag is set

input error flag is reset

output error flag is reset

conflict for control of interface flag is reset

computer in process of becoming controller in charge flag is reset

computer in process of acknowledging an unidentified command from

another device flag is reset

computer in process of entering the controller active state following a 	TRANSFER command flag is reset

REN remote enable remote message being sent flag is reset

non fatal error flag is reset

Remote Messages

IFC

IFC false

Description

The primary address of the computer, as specified by the BBC DEVICE NO

command, is set before sending the IFC interface clear remote message. This

places the interface in a known quiescent state. The computer exits the

command as controller in charge.

Associated Commands

BBC DEVICE NO

- 23 -

DEVICE CLEAR

Purpose

To clear all devices

Example

100 PRINT#cmd%,"DEVICE CLEAR"

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

No change

Remote Messages

ATN

true	DCL

Description

After this command has been executed all devices with a device clear

capability are initialised.

Associated Commands

SELECTED DEVICE CLEAR

- 24 -

END OF STRING

Purpose

To specify the characters delimiting device dependent messages.

Example

100 endchr$=CHR$(13)+CHR$(10)

110 PRINT#cmd%,"END OF STRING",endchr$

Notes

String constants and variables of 1 or 2 characters may be used to specify

the end of string characters. The default character is LF (line feed).

IEEE Status Required before Execution

No requirement

IEEE Status after Execution

No change

Remote Messages

None

Description

The computer stores the End of String characters and sends them after each

string of data bytes it outputs and expects to receive them after each

string of data bytes it inputs.

- 25 -

GOTO LOCAL

Purpose

To return the referenced devices to local control.

Example

100 voltmeter%=OPENIN("9")

110 plotter%=OPENIN("5")

120 PRINT#cmd%,"GO TO LOCAL",voltmeter%,plotter%,"EXECUTE"

Notes

The last channel must be followed by the 'EXECUTE' command.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device or devices have been LISTENed flag is reset.

Remote Messages

ATN

true	UNIL

true	MLA 1

true	MSA 1	if specified

true	MLA 2

true	MSA 2	if specified

etc

true	MLA n

true	MSA n	if specified

true	GTL

true	UNL

Description

After this command has been executed, devices with a remote capability

referenced by the previously OPENINed channels are returned to local

control.

Associated Commands

REMOTE ENABLE

REMOTE DISABLE

LISTEN

LOCAL LOCKOUT

- 26 -

LISTEN

Purpose

To address the referenced devices as listeners.

Example

100 voltmeter%=OPENIN("2,1")

110 printer%=OPENIN("5")

120 PRINT#cmd%,"LISTEN",voltmeter%,printer%,"EXECUTE"

Notes

Up to 14 devices may be LISTENed at any one time. The last channel must be

followed by the 'EXECUTE' command.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

controller in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status After Execution

The	device or devices have been LISTENed flag is set.

Remote Messages

ATN

true	MLA 1

true	MSA 1	if specified

true	MLA 2

true	MSA 2	if specified

etc

true	MLA n

true	MSA n	if specified

Description

The computer sends the MLA, my listen address, and MSA, my secondary address

if specified, remote messages of the devices referenced by the previously

OPENINed channels.

Associated Commands

UNLISTEN

- 27 -

LOCAL LOCKOUT

Purpose

To prevent conflict between the remote and local operation of all devices.

Example

100 PRINT#cmd%,"LOCAL LOCKOUT"

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

No change

Remote Messages

ATN

True	LL0

Description

After this command has been executed, devices with a local lockout

capability may only be returned to local (front panel) from remote

(interface) control by using the GO TO LOCAL command.

Associated Commands

REMOTE ENABLE

REMOTE DISABLE LISTEN

GO TO LOCAL

- 28 -

PARALLEL POLL DISABLE

Purpose

To disable the referenced devices so that they are unable to take part in a

parallel poll.

Example

100 analyser%=OPENIN("11")

110 voltmeter%=OPENIN("6")

120 PRINT#cmd%,"PARALLEL POLL DISABLE",analyser%,voltmeter%,"EXECUTE"

Notes

The last channel must be followed by the 'EXECUTE' command.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device or devices have been LISTENed flag is reset.

Remote Messages

ATN

true	UNL

true	MLA 1

true	MSA 1	if specified

true	MLA 2

true	MSA 2	if specified

etc

true	MLA n

true	MSA n	if specified

true	PPC

true	PPD

true	UNL

Associated Commands

PARALLEL POLL ENABLE

PARALLEL POLL UNCONFIGURE

PARALLEL POLL REQUEST

- 29 -

PARALLEL POLL ENABLE

Purpose

To enable the referenced device to take part in a parallel poll.

Example

100 spectrumanalyser%=OPENIN("9")

110 PRINT#cmd%,"PARALLEL POLL ENABLE",spectrumanalyser%,10

Notes

The parallel poll response remote message is specified by the least

significant 4 bits of an integer constant or variable.

Bit 3 specifies the sense of the response and Bits 0-2 specify the response.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device or devices have been LISTENed flag is reset

Remote Messages

ATN

true	UNL

true	MLA

true	MSA

true	PPC

true	PPE

true	UNL

Associated Commands

PARALLEL POLL DISABLE

PARALLEL POLL UNCONFIGURE

PARALLEL POLL REQUEST

- 30 -

PARALLEL POLL REQUEST

Purpose

To conduct a parallel poll.

Example

100 PRINT#cmd%,"PARALLEL POLL REQUEST"

110 INPUT#cmd%,pollresponse%

Notes

The response may only be returned to the least significant byte of an

integer variable via the command channel.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

CLEAR command executed flag must be set

IEEE Status after Execution

No change

Remote Messages

ATN	IDY

true	true	PPR

Description

If a TRANSFER is in progress then the computer waits until the current data

byte has been transferred before conducting the PARALLEL POLL; it then

restarts the transfer. In this way no data is lost.

If no TRANSFER is in progress the PARALLEL POLL is conducted immediately.

Associated Commands

PARALLEL POLL ENABLE

PARALLEL POLL DISABLE

PARALLEL POLL UNCONFIGURE

- 31 -

PARALLEL POLL UNCONFIGURE

Purpose

To render all devices unable to take part in a parallel poll.

Example

100 PRINT#cmd%,"PARALLEL POLL UNCONFIGURE"

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

No change

Remote Messages

ATN

true	PPU

Associated Commands

PARALLEL POLL ENABLE

PARALLEL POLL DISABLE

PARALLEL POLL REQUEST

- 32 -

READ BINARY

Purpose

To enable the input of a specified number of data bytes.

Example

90 DIM reading%(511)

100 recorder%=OPENIN("5")

110 PRINT#cmd%,"TALK",recorder%

120 PRINT#cmd%,"READ BINARY",512

130 FORI%=OT0511

140 reading%(I%)=BGET#data%

150 NEXT

160 CLOSE#recorder%

Notes

Integer constants and variables in the range 1 to &7FFFFFF may be used for

the number of data bytes.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

a device has been TALKed flag must be set

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	binary coded data being handled flag is set.

Remote Messages

ATN

false	DAB 1		data bytes received

false	DAB 2

etc

false	DAB n and END	the END remote message is not required

true

Description

This command enables the input of long, in excess of 255 data bytes, binary

coded device dependent remote messages. The computer flags binary coded data

being handled in the IEEE status word and addresses itself as a listener. A

single data byte is input in response to each BGET#. When the specified

number of data bytes have been input, the computer resets the flag in the

IEEE status word and unaddresses itself as a listener. If the END end remote

message is received during a READ BINARY command, this is flagged in the

IEEE status word. However, it does not terminate the input.

Associated Commands

TALK

UNTALK

- 33 -

REMOTE DISABLE

Purpose

To return all devices to local control.

Example

100 PRINT#cmd%,"REMOTE DISABLE"

IEEE Status Required before Execution

The	CLEAR command executed flag must be set.

IEEE Status after Execution

The	REN remote enable remote message being sent flag is reset.

Remote Messages

REN	false

Description

After this command has been executed all devices with a remote local

capability are returned to local control, including those put in the local

lockout state by a previous LOCAL LOCKOUT command.

Associated Commands

REMOTE ENABLE

LISTEN

LOCAL LOCKOUT

GO TO LOCAL

- 34 -

REMOTE ENABLE

Purpose

To enable the remote operation of devices.

Example

100 PRINT#cmd%,"REMOTE ENABLE"

IEEE Status Required before Execution

The	CLEAR command executed flag must be set.

IEEE Status after Execution

The	REN remote enable message being sent flag is set.

Remote Messages

REN

Description

After this command has been executed, the computer can change devices from

local (front panel) control to remote (Interface) control using the LISTEN

command.

Associated Commands

REMOTE DISABLE LISTEN

LOCAL LOCKOUT

GO TO LOCAL

- 35 -

REQUEST CONTROL

Purpose

To request that control be transferred back to the computer.

Example

100 PRINT#cmd%,"REQUEST CONTROL"

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

CLEAR command executed flag must be set

computer in process of becoming controller in charge flag must be unset

IEEE Status after Execution

No change

Remote Messages

ATN				sent by current controller in charge

SRQ

STB, RQS and SRQ false	when polled by controller in charge

ATN				sent by computer

true				after TCT received from current controller in

charge

Description

The computer sends the SRQ service request remote message and then exits the

command. When the controller in charge conducts a serial poll the STB status

byte and RQS request service remote messages are sent. The seven bits of the

STB remote message are all 0. If the controller in charge sends the TCT,

take control, remote message the computer becomes controller in charge and

sets the appropriate flag in the IEEE status word.

Associated Commands

TAKE CONTROL

CLEAR

- 36 -

SELECTED DEVICE CLEAR

Purpose

To clear the referenced devices.

Example

100 plotter%=OPENIN("7")

110 counter%=OPENIN("3")

120 PRINT#cmd%,"SELECTED DEVICE CLEAR",plotter%,counter%,"EXECUTE"

Notes

The last channel must be followed by the 'EXECUTE' command.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device or devices have been LISTENed flag is reset.

Remote Messages

ATN

true	UNL

true	MLA 1

true	MSA 1 if specified

true	MLA 2

true	MSA 2 if specified

etc

true	MLA n

true	MSA n if specified

true	SDC

true	UNL

Description

After this command has been executed the devices with the device clear

capability referenced by the previously OPENINed channels are initialised.

Associated Commands

DEVICE CLEAR

- 37 -

SERIAL POLL

Purpose

To conduct a serial poll of the referenced device.

Example

100 logicanalyser%=OPENIN("7")

110 PRINT#cmd%,"SERIAL POLL",Iogicanalyser%,1

120 INPUT#cmd%,statusmessage$

Notes

Integer constants and variables in the range 1 to 255 may be used for the

number of status bytes sent by the device to be polled. The status bytes may

only be returned to a string variable via the command channel.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

controller in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

If the RQS, request service, remote message is received, the SRQ, service

request, remote message received flag is reset.

The	device has been TALKed flag is reset

The	device or devices have been LISTENed flag is reset

Remote Messages

ATN

true	UNL

true	SPE

true	MTA

true	MSA		if specified

false	STB 1 and RQS?

false	STB 2

etc

false	STB n

true	SPD

true	UNT

Description

Following the receipt of an SRQ, service request, remote message the SERIAL

Poll command should be used to determine which device requires service.

Associated Commands

STATUS

- 38 -

STATUS

Purpose

To return information on the current state of the interface.

Example

100 PRINT#cmd%,"STATUS"

120 INPUT#cmd%,state%

Notes

The status may only be returned to an integer variable via the command

channel.

IEEE Status Required before Execution

No requirement

IEEE Status after Execution

No change

Remote Messages

None

Description

The individual bits of the status word indicate the following

- 39 -

Bit

0 timeout on output

1 timeout on input

2 computer is controller in charge

3 TRANSFER in progress

4 string of more than 255 characters received

5 SRQ service request remote message received

6 END end remote message received

7 Incomplete source handshake

8 binary coded data being handled

9 TIMEOUT ON

10 "COMMAND" file open

11 "DATA" file open

12 a device has been TALKed

13 a device or devices have been LISTENed

14 BBC DEVICE NO executed

15 CLEAR executed

16 input error

17 output error

18 conflict for control of interface

19 computer in process of becoming controller in charge

20 computer in process of acknowledging an unidentified command from another

device

21 computer in process of entering the controller active state following a

TRANSFER command

22 unused

23 REN remote enable remote message being sent

24 unused

25 unused

26 unused

27 unused

28 unused

29 unused

30 unused

31 nonfatal error

- 40 -

TAKE CONTROL

Purpose

To pass control to another device.

Example

100 secondcontroller%=OPENIN("2')

110 PRINT#cmd%,"TAKE CONTROL",secondcontroller%

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	computer is controller in charge flag is reset

device or devices have been LISTENed flag is reset

device has been TALKed flag is reset

Remote Messages

ATN

true	UNL

true	MTA

true	MSA if specified

true	TCT

false

Description

After this command has been executed the computer may regain control of the

interface through REQUEST CONTROL or CLEAR commands.

Associated Commands

REQUEST CONTROL

CLEAR

- 41 -

TALK

Purpose

To address the referenced device as a talker.

Example

100 digitalvoltmeter%=OPENIN("2,1")

110 PRINT#cmd%,"TALK",digitalvoltmeter%

Notes

Only one device may be TALKed at any one time.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

controller in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device has been TALKed flag is set.

Remote Messages

ATN

true	MTA

true	MSA if specified

Description

The computer sends the MTA, my talk address, and, if specified, the MSA, my

secondary address, remote messages of the device referenced by the

previously OPENINed channel.

Associated Commands

UNTALK

- 42 -

TIMEOUT OFF

Purpose

To disable timeouts.

Example

100 PRIN#cmd%,"TIMEOUT OFF"

Notes

Timeouts are not a part of the IEEE488 standard.

TIMEOUT OFF is the default IEEE status.

IEEE Status Required before Execution

No requirement.

IEEE Status after Execution

The	TIMEOUT ON flag is reset.

Remote Messages

None

Description

With TIMEOUT OFF, if a device fails to respond during a source or acceptor

handshake, then the ESC key must be pressed to exit from a command.

Associated Commands

TIMEOUT ON

- 43 -

TIMEOUT ON

Purpose

To enable timeouts.

Example

100 PRINT#cmd%,"TIMEOUT ON"

Notes

Timeouts are not a part of the IEEE488 standard.

TIMEOUT OFF is the default IEEE status.

IEEE Status Required before Execution

No requirement.

IEEE Status after Execution

The	TIMEOUT ON flag is set.

Remote Messages

None

Description

After this command has been executed and a device fails to respond in 2.55

seconds during a source or acceptor handshake then the appropriate timeout

error is flagged in the IEEE status word.

Associated Commands

TIMEOUT OFF

- 44 -

TRANSFER

Purpose

To initiate the transfer of data bytes between an active talker and all

active listeners.

Example (a)

100 printer%=OPENIN("4,2")

110 counter%=OPENIN("9")

120 PRINT#cmd%,"LISTEN",printer%,"EXECUTE"

130 PRINT#cmd%,"TALK",counter%

140 PRINT#cmd%,"TRANSFER",1024

150 PRINT#cmd%,"UNTALK"

160 PRINT#cmd%,"UNLISTEN"

Example (b)

100 printer%=OPENIN("4,2")

110 counter%=OPENIN("9")

120 PRINT#cmd%,"LISTEN",printer%,"EXECUTE"

130 PRINT#cmd%,"TALK",counter%

140 PRINT#cmd%,"TRANSFER",0

150

.

.

.

190 REPEAT

200 PRINT#cmd%,"STATUS"

210 INPUT#cmd%,status488%

220 UNTIL (status488% AND 8) = 8

Notes

Integer constants and variables in the range 0 to &7FFFFFF may be used for

the number of data bytes to be transferred.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

a device has been TALKed flag must be set

a device or devices have been LISTENed flag must be set

conflict for control of interface flag must be unset

controller in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

When the number of data bytes is specified as 1 to 134217727, example (a),

the END end remote message received flag is set if this occurs during the

TRANSFER. When the number of data bytes is specified as 0, example (b). the

TRANSFER in progress flag remains set until the END end remote message is

received, it is then reset.

Note that the END end remote message received flag is not set.

- 45 -

Remote Messages

ATN

false DAB 1		data bytes disregarded by computer

false DAB 2

etc

false DAB n and END	the END remote message is not required if the 	true				number of data bytes is not 0

Description

The TRANSFER command can be used in two ways:

Example (a)

The computer initiates the TRANSFER and after the talker has sent the

specified number of data bytes exits the command.

Example (b)

The computer initiates the TRANSFER and exits the command. When the talker

sends the END end remote message with the last data byte transferred, the

computer resets the TRANSFER in progress flag in the IEEE status word. Thus,

the computer can undertake other tasks while the TRANSFER is in progress and

detect its completion by reading the IEEE status word.

Associated Commands

TALK

UNTALK

LISTEN

UNLISTEN

- 46 -

TRIGGER

Purpose

To initiate simultaneous operation of the referenced devices.

Example

100 generator%=OPENIN("9")

110 voltmeter%=OPENIN("14")

120 PRINT#cmd%,"TRIGGER",generator%,voltmeter%,"EXECUTE"

Notes

The last channel must be followed by the 'EXECUTE' command.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device or devices have been LISTENed flag is reset.

Remote Messages

ATN

true	UNL

true	MLA 1

true	MSA 1 if specified

true	MLA 2

true	MSA 2 If specified

etc

true	MLA n

true	MSA n if specified

true	GET

true	UNL

Description

This command causes all devices with the device trigger capability

referenced by the previously OPENINed channels to start some basic

operation.

Associated Commands

None

- 47 -

UNLISTEN

Purpose

To unaddress all active listeners.

Example

100 PRINT#cmd%,"UNLISTEN"

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

controller in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device or devices have been LISTENed flag is reset.

Remote Messages

ATN

true UNL

Description

The computer sends the UNL unlisten remote message to all active listeners.

Associated Commands

LISTEN

- 48 -

UNTALK

Purpose

To unaddress an active talker.

Example

100 PRINT#cmd%,"UNTALK"

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

controller in process of acknowledging an unidentified command from

another device flag must be unset

IEEE Status after Execution

The	device has been TALKed flag is reset.

Remote Messages

ATN

true UNT

Description

The computer sends the UNT untalk remote message to the active talker.

Associated Commands

TALK

- 49 -

WRITE BINARY

Purpose

To enable the output of a specified number of data bytes.

Example

100 analyser%=OPENIN("11")

110 PRINT#cmd%="LISTEN",analyser%,"EXECUTE"

120 PRINT#cmd%,"WRITE BINARY",1024

130 FOR 1% = 1 TO 1024

140 BPUT#data%,waveform%(I%)

150 NEXT

160 CLOSE#analyser%

Notes

Integer constants and variables in the range 1 to &7FFFFFF may be used for

the number of data bytes.

IEEE Status Required before Execution

The	computer is controller in charge flag must be set

TRANSFER in progress flag must be unset

a device or devices have been LISTENed flag must be set

CLEAR command executed flag must be set

conflict for control of interface flag must be unset

computer in process of acknowledging unidentified command from another

device flag must be unset

IEEE Status after Execution

The	binary coded data being handled flag is set

Remote Messages

ATN

true	UNT

false	DAB 1		data bytes sent

false	DAB 2

etc

false	DAB n and END

true

Description

This command enables the output of long, over 255 data bytes, binary coded

device dependent messages. The computer flags binary coded data is being

handled in the IEEE status word and after sending the UNT untalk remote

message addresses itself as a talker. A single data byte is output in

response to each BPUT#. The END end remote message is sent with the last

data byte. When the specified number of data bytes have been output the

computer resets the flag in the IEEE status word and unaddresses itself as a

talker.

Associated Commands

LISTEN

UNLISTEN

- 50 -

7 USING THE IEEE FILING SYSTEM IN ASSEMBLER

Read Chapter 43 of the BBC Microcomputer User Guide before writing assembler

programs that use the IEEE488 interface. General information on using a

filing system in assembler is presented there. This chapter gives details of

the relevant operating system calls when using the IEEEFS.

OSFIND	&FFCE	Open or close a channel

OSGBPB	&FFD1	Not supported by IEEEFS

OSBPUT	&FFD4	Send a single byte to channel from A

OSBGET	&FFD7	Receive a single byte from channel in A

OSARGS	&FFDA	Not fully supported by IEEEFS

OSFILE	&FFDD	Not supported by IEEEFS

OSWORD	&FFF1	Perform IEEE commands using parameter block to pass parameters

OSCLI	&FFF7	Interpret IEEEFS command line

Throughout this section the term file is used to maintain compatibility with

Chapter 43 of the BBC Microcomputer User Guide. However, it should be noted

that IEEEFS 'files' are not true files in that they have no storage

capability and serve only to pass information to the IEEFS.

OSFIND

Opens a file for reading and writing. The routine is entered at &FFCE which

indirects through &21C. The value in A determines the type of operation.

A =&00 causes a file or files to be closed

A<>&00 causes a file to be opened

If A<>0 then Y (high byte) and X (low byte) must contain the address of a

location in memory which contains the file name terminated with CR (0D).

On exit A will contain the channel number allocated to the file for all

future operations. If A = 0 then the IEEEFS was unable to open the file.

If A=0 then a file, or all files will be closed depending on the value of Y.

Y=0 will close all files, otherwise the file whose channel number is in Y

will be closed.

On exit C, N, V and Z are undefined and D=0. The interrupt state is

preserved.

- 51 -

OSBPUT

This call is used to send data. It write (puts) a byte in A to the file

previously opened using OSFIND. The routine is entered at &FFD4 which

indirects through &218. On entry Y contains the channel number allocated by

OSFIND.

On exit A, X and Y are preserved, N, V and Z are undefined and D=0. The

interrupt state is preserved.

OSBGET

This call is used to receive data. It gets (reads) a byte from a file into

A. The file must have been previously opened using OSFIND and the channel

number allocated must be in V. The routine is entered at &FFD7 which

indirects through &216.

On exit C=0 indicates a valid character in A. X and Y are preserved, N V and

Z are undefined and D=0. The interrupt state is preserved.

OSARGS

This routine is entered at &FFDA and indirects through &214. The IEEEFS only

supports a call with A=&00 and Y=&00; this returns with A=&07 indicating

that the IEEEFS is in use.

On exit X and Y are preserved, C, N, V and Z are undefined and D=0. The

interrupt state is preserved.

OSWORD

This routine is entered at &FFF1 and indirects through &20C. IEEE commands

are sent and information on the state of the interface received through the

use of a single OSWORD call with A=&80. This fulfils the same function as

the COMMAND channel. Y (high byte) and X (low byte) point to a parameter

block. The structure of this is:

first byte       : no of bytes in argument

second byte      : no of bytes in result

third byte       : command code

subsequent bytes : as defined by the first three bytes

Details of each command are given in the table.

- 52 -

OSCLI

This call is used to select the IEEEFS by placing the address of the command

line IEEE in Y (high byte) and X (low byte) and calling &FFF7 which

indirects through &2xx. The command line should not start with an asterisk

and must end with &0D. Any leading spaces will be stripped.

DATA

When data is sent using the OSBPUT and OSBGET calls, strings are transferred

as type (&00) length and characters (in reverse order). When binary coded

data is being handled it is sent and received a byte at a time.

Example

Typical sequence of operating system calls.

Select IEEFFS by calling OSCII

Open COMMAND and DATA channels by using OSFIND

Execute IEEE commands eg CLEAR, REMOTE ENABLE etc by calling OSWORD

Open 'device' channels by calling OSFIND

Listen or talk devices by calling OSWORD

Send and receive data by calling OSBPUT and OSBGET

Close channels by calling OSFIND

- 53 -

Structure of OSWORD Block

XY+   0    1     2       3    4     5     .    .

Command                   arg0 arg1 command   b0    b1    b2    .    .

BBC DEVICE NO               1    0    &40   dev no

CLEAR                       0    0    &50

DEVICE CLEAR                0    0    &62

END OF STRING            1 or 2  0    &81    eos2  eos1

GO TO LOCAL                 n    0    &22    chn    .     .    ch2  ch1

LISTEN                      n    0    &20    chn    .     .    ch2  ch1

LOCAL LOCKOUT               0    0    &61

PARALLEL POLL DISABLE       n    0    &21    chn    .     .    ch2  ch1

PARALLEL POLL ENABLE        2    0    &48    ch1   ppr

PARALLEL POLL REQUEST       0    1    &02    ppr

PARALLEL POLL UNCONFIGFURE  0    0    &60

READ BINARY                 4    0    &14    cnt3  cnt2  cnt1  cnt0

REMOTE DISABLE              0    0    &71

REMOTE ENABLE               0    0    &70

REQUEST CONTROL             0    0    &64

SELECTED DEVICE CLEAR       n    0    &24    chn    .     .    ch2  ch1

SERIAL POLL                 2    n    &44    ch1  pl1cnt  bn    .    b1

STATUS                      0    4    &01     b3    b2    b1    b0

TAKE CONTROL                1    0    &42    ch1

TALK                        1    0    &41    ch1

TIMEOUT OFF                 0    0    &52

TIMEOUT ON                  0    0    &51

TRANSFER                    4    0    &10    cnt3  cnt2  cnt1  cnt0

TRIGGER                     n    0    &28    chn    .     .    ch2  ch1

UNLISTEN                    0    0    &54

UNTALK                      0    0    &58

WRITE BINARY                4    0    &18    cnt3  cnt2  cnt1  cnt0

NOTE:

The bytes passed in XY+0 and XY+1 must be the length of the whole control

block, including the length bytes. Consequently, the above arg0 and arg1

values must be increased by 2 before being placed in the control block.

- 54 -

8 HARDWARE

The interface uses the Texas Instruments TMS 9914A GPIB Adapter, together

with the SN75160A and SN75162A octal GPIB transceivers. The TMS 9914A

interfaces the computer via the 1MHz bus to the IEEE488 bus. It relieves the

computer's processor of the task of handling IEEE488 protocol when commands

and data are sent and received. The TMS 9914A generates interrupts and so

the processor does not have to spend time polling it.

As supplied, the hardware is configured as system controller, Molex link S1

on interface PCB in position A, and with the IEEE488 data bus to be driven

by open collector outputs, Molex link S2 on interface PCB in position C.

A 5 MHz crystal provides a clock for timing.

The TMS 9914A is connected to the 1MHz bus in accordance with the guidelines

given in Ref 2. The interface is supplied with four terminating resistor

packs. They should be removed if the interface is not to be the last device

on the 1MHz bus.

The read and write registers of the interface are located at addresses &FC20

to &FC27:

&FC20 : Interupt Status 0    / Interupt Mask 0

&FC21 : Interupt Status 1    / Interupt Mask 1

&FC22 : Address Status       /        -

&FC23 : Bus Status           / Auxilary Control

&FC24 :           -          / Address

&FC25 :           -          / Serial Poll

&FC26 : Command Pass Through / Parallel Poll

&FC27 : Data In              / Data Out

The user who wishes to program the TMS 9914A directly in order to configure

the computer as a simple talker-listener should obtain Ref 3, and change the

Molex link S1 on the interface PCB from position A to position B.

- 55 -

REFERENCES

Ref 1	IEEE Std 488-1978

IEEE488 Standard Digital Interface for Programmable Instrumentation

The Institute of Electrical and Electronic Engineers Inc 1978

Ref 2	Application Note No 1

BBC Microcomputer 1MHz Bus

Acorn Computers Ltd 1982

Ref 3	TMS 9914 GPIB Adapter Data Manual

Texas Instruments

- 56 -

APPENDIX 1

IEEE Commands                        Minimum Abbreviation

BBC DEVICE NO                        B.

CLEAR                                C.

DEVICE CLEAR                         D.

END OF STRING                        E.

EXECUTE                              E.

GO TO LOCAL                          G.

LISTEN                               L.

LOCAL LOCKOUT                        LO.

PARALLEL POLL DISABLE                PARALLEL POLL D.

PARALLEL POLL ENABLE                 PARALLEL POLL E.

PARALLEL POLL REQUEST                P.

PARALLEL POLL UNCONFIGURE            PARALLEL POLL U.

READ BINARY                          R.

REMOTE DISABLE REMOTE                D.

REMOTE ENABLE                        REM.

REQUEST CONTROL                      REQ.

SELECTED DEVICE CLEAR                SEL.

SERIAL POLL                          SE.

STATUS                               S.

TAKE CONTROL                         TAK.

TALK                                 T.

TIMEOUT OFF                          TIMEOUT OFF

TIMEOUT ON                           TI.

TRANSFER                             TR.

TRIGGER                              TRI.

UNLISTEN                             UNI.

UNTALK                               U.

WRITE BINARY                         W.

Note - Standard IEEE488 mnemonics can be used as string variable names.

Example

100 PPU$="PARALLEL POLL UNCONFIGURE"

110 PPE$="PARALLEL POLL ENABLE"

120 PPR$="PARALLEL POLL REQUEST"

130 PPD$="PARALLEL POLL DISABLE"

.

.

.

650 spectrumanalyser%=OPENIN("9")

660 PRINT#cmd%,PPU$

670 PRINT#cmd%,PPE$,spectrumanalyser%,10

.

.

.

940 PRINT#cmd%,PPR$

950 INPUT#cmd%,response%

.

.

990 PRINT#cmd%,PPD$,spectrumanalyser%,"EXECUTE"

- 57 -

APPENDIX 2

IEEEFS ERRORS

Error number

&DC syntax

&88 out of range

&DE channel

&8A no device listened type mismatch

&8B no BBC DEVICE NO

&8D string too long

&89 no device talked

&8C CLEAR not executed

&FE Bad command

- 58 -

APPENDIX 3

INTERFACE REMOTE MESSAGES

List of Messages Referenced in this Guide

Mnemonic

ATN	attention

DAB	data byte

DCL	device clear

END	end

EOS	end of string

GET	group execute trigger

GTL	go to local

IDY	identify

IFC	interface clear

LLO	local lockout

MIA	my listen address

MTA	my talk address

MSA	my secondary address

PPC	parallel poll configure

PPE	parallel poll enable

PPD	parallel poll disable

PPR	parallel poll response

PPU	parallel poll unconfigure

REN	remote enable

RQS	request service

SDC	selected device clear

SPD	serial poll disable

SPE	serial poll enable

SRQ	service request

STB	status byte

TCT	take control

UNL	unlisten

UNT	untalk

- 59 -

APPENDIX 4

INTERFACE SPECIFICATION - note 1

SH		source handshake function

SH1		complete capability

AH		acceptor handshake function

AH1		complete capability

T		talker function - note 2

T5		complete capability

L		listener function - note 2

L3		complete capability

SR		service request

SR1		complete capability

RL		remote local function

RL0		no capability

PP		parallel poll function

PP0		no capability

DC		device clear function

DC0		no capability

DT		device trigger function

DT0		no capability

C		control function

C1, C2, C3	complete capability

C4 and C5

Note 1 - This specification applies to the IEEEFS and interface hardware

configured as system controller.

Note 2 - The interface only sends and receives data when the computer

addresses itself to talk or listen, not when addressed by other devices.

Power supply 220-24OVAC 50-60Hz 3 watt

Manufactured in accordance with BS415/79 Class 1

Designed and manufactured to comply with Acorn Computers Ltd 1MHz Bus

Peripheral Specification.

- 60 -

APPENDIX 5

FITTING THE IEEEFS EPROM

CAUTION - STATIC SENSITIVE DEVICES

DO NOT HANDLE UNTIL YOU HAVE READ THESE INSTRUCTIONS

Please read these instructions carefully before attempting this conversion.

If you are not completely confident of your ability to carry out this

conversion yourself it will be in your own interests to take this unit to an

authorised BBC Microcomputer dealer for him to carry out this upgrade.

1	Determine which operating system is fitted in your computer by typing

*FX0 and noting the version number displayed

2	Switch off and disconnect the computer from the mains supply

WARNING: DO NOT REMOVE THE LID OF THE COMPUTER WHILST IT IS CONNECTED

TO THE MAINS

3	Unscrew the two large-headed screws on the rear panel of the computer

and the two large-headed screws on the underside of the computer (near

the front). Carefully remove the lid. 4 Unscrew the two (or three on

some computers) bolts securing the keyboard assembly to the lower case

and CAREFULLY unplug the 17-way ribbon cable connecting the keyboard

assembly to the main Printed Circuit Board (PCB) by pulling the socket

away from PL13 on the main PCB.

5	Move the keyboard away from the main PCB until the loudspeaker plug is

accessible. Unplug the loudspeaker by CAREFULLY removing the socket from

PL15. Place the keyboard assembly to one side on a flat surface.

6	If your computer contains an operating system of version 1.00 or later

proceed to instruction 9

7	The current Machine Operating System (MOS) is situated in the bottom

right-hand corner of the main PCB in one of two formats as follows:

a)	One integrated circuit (IC) in socket 1C51 identified by the legend

"B02" printed on top. If you have this format carefully remove the MOS

IC from 1C51 using an IC extractor (a small screwdriver used with

extreme care will do).

b)	Four ICs in sockets IC52, IC88, IC100 and IC101 (these sockets are all

next to each other despite the numbering) identified with the legends

"C", "D", "E" and "F" respectively printed on the top. If this is the

case the BASIC language IC (identifiable by the legend "B01" printed on

the top) and all four MOS ICs should be removed as described in 'a'

above. Replace the BASIC IC in socket IC100 as described in '8'.

8	The new MOS is supplied in a single IC identified by the legend "B04"

printed on the top. Insert the ROM carefully into the socket 1C51

ensuring that all the 'legs' locate properly into the holes and are not

'bent under' upon insertion. The IC should be orientated so that the

semi-circular cutout in the top of the IC is at the top, i.e. away from

the keyboard. Avoid touching the IC legs with your fingers wherever

possible.

- 61 -

Now set the following links:

S18 North

S19 East

S20 North

S21 Two off East/West

S22 North

S32 West

S33 West

N.B. North is the edge of the PCB nearest to the back of the case and

South is the edge nearest to the keyboard. East and West follow

logically from these.

9	Insert the IEEEFS EPROM IC into any free socket (IC88, IC100 or IC101)

as described in 8. If you have a disc filing system (DFS) ensure that

the IEEEFS EPROM is inserted to the right of the DFS IC.

10	Reassemble the BBC Microcomputer by reversing instructions 2.

11	To check that the IEEEFS EPROM has been fitted correctly type

*HELP IEEE

The computer should display

IEEE Filing System <version number>

OS <version number>

If it does not, disconnect the computer from the mains supply and ensure

that all the instructions have been followed correctly. Reconnect to

the supply and try again. If you still have problems contact an

authorised dealer.

- 62 -

APPENDIX 6

REMOVAL OF TERMINATING RESISTOR PACKS

The interface is supplied with four terminating resistor packs. They should

be removed following the instructions below if the interface is not to be

the last device on the 1MHz bus.

1	Switch off the interface and disconnect it from the mains supply

WARNING - DO NOT REMOVE THE LID OF THE INTERFACE WHILST IT 15 CONNECTED

TO THE MAINS

2	Disconnect the 1MHz Bus connecting cable from the computer or other 1MHz

bus peripheral.

3	The two halves of the case may be separated by undoing three fixing

screws two on the rear panel and one on the underside in the centre at

the front. When reassembling push the two halves together whilst

tightening the two rear fixing screws. Take care not to lose the two

captive nuts for the rear fixing screws which are pushed onto the upper

moulding.

4	The four resistor packs, RP1, RP2, RP3 and RP4, are located on the PCB

towards the front of the case. Note the orientation of the resistor

packs for future reference. Carefully remove them using an IC extractor

(a small screwdriver used with extreme care will do) and store them

safely so that they can be refitted in the future should the interface

become the last peripheral on the 1MHz bus.

5	Reassemble by following steps 2 and 3 in the reverse order.

- 63 -

INDEX

Abbreviations          Appendix 1		OSARGS                         52

Assembler                      51		OSBGET                         52

OSBPUT                         52

BBC DEVICE NO       5, 17, 18, 22		OSCLI                          52

BGET#                       9, 20		OSFIND                         51

BPUT#                       9, 20		OSWORD                     52, 54

CLEAR       5, 14, 15, 17, 21, 23		PARALLEL POLL                  12

CLOSE                           4			-DISABLE             12, 29

COMMAND channel             4, 17			-ENABLE              12, 30

Controller in charge           15			-REQUEST             12, 31

-UNCONFIGURE         12, 32

DATA channel                4, 17		Primary address             4, 18

Data input           6, 9, 18, 53		PRINT#                          6

Data output          6, 9, 19, 53

Delimiters                      5		READ BINARY                 9, 33

DEVICE CLEAR               14, 24		REMOTE DISABLE             13, 34

REMOTE ENABLE           5, 13, 35

END message                     7		REQUEST CONTROL            15, 36

END OF STRING               5, 25

EOI                             7		Secondary address           4, 18

EPROM - fitting        Appendix 5		SELECTED DEVICE CLEAR      14, 37

EXECUTE                         6		SERIAL POLL                11, 38

Source handshake                7

GO TO LOCAL                13, 26		STATUS                  7, 39, 40

Group execute trigger      14, 47		System controller              15

Hardware                       55		TAKE CONTROL               15, 41

TALK                        6, 42

IEEE488 - standard              1		Terminating resistors  Appendix 6

- *IEEE               3, 17		TIMEOUT OFF                13, 43

- IEEEFS              3, 51		TIMEOUT ON                 13, 44

INPUT#                          6		TRANSFER               10, 45, 46

TRIGGER                    14, 47

LISTEN                      6, 27

LOCAL LOCKOUT              13, 28		UNLISTEN                    6, 48

UNTALK                      6, 49

OPENIN                          4

OPENOUT                         4		WRITE BINARY                9, 50

- 64 -

ACORN COMPUTERS LTD

Fulbourn Road

Cherry Hinton

Cambridge

CB1 4JN